# CERTIFICATE OF EC TYPE EXAMINATION

# NP635.64D-4E1RLUI and NP635.66D-4E1RlnI

Issued to

In respect of

A three-phase electricity meter type family.

In accordance with

The Swedish Act on Metrology and Verification STAFS 2006:4 and STAFS 2006:7 dated  2006-08-21, transposing Directive 2004/22/EC on measuring instruments (MID).

Applicable essential requirements

Annex 1

Annex MI-003, ACTIVE ELECTRICAL METERS

Harmonised standards and normative documents used

- SS EN 50470-1:2005, SS EN 50470-3:2005

Further applied documents

- WELMEC 7.2, Software Guide (Issue 1)

Rated operating conditions

| Measurand: | Energy/kWh | Mechanic influence class: | M1 |
| --- | --- | --- | --- |
| Accuracy class: | B | Climatic influence class: | -40 to +70C |
| Measurement range: | 999999,99 kWh (configurable parameter) | Electromagnetic influence class: | E2 |

The meter is also tested and approved in accordance with the following standards and methods:

- SS EN 62053-23, Reactive energy metering, class 2.

- SP-Method 1618, Impulse voltage withstand of 12 kV, test modified from SS EN 50470-1.

Valid until

15th December, 2016.

The principal characteristics, approval conditions are set out in the appendix hereto, which forms part of the approval documents and consists of 4 pages. All the plans, schematic diagrams and documentations are recorded under reference MTeP6070822.

Borås xxth January 2007

SP Swedish National Testing and Research Institute

Certification

Lennart Månsson	Stefan Svensson

Certification Manager   	Certification Officer

The instruments / measuring systems must correspond with the following specifications:

1	Design of the instrument

1.1	Construction

The instrument is designed to measure electrical energy at 230 V, 50 Hz in a three-phase four-wire installation. It communicates via power line communication, proprietary of the manufacturer. It may have or not a breaker capable to disconnect the load on command over the communication device.

1.2	Measurand sensor

The current sensors are transformers capable to handle high levels of DC current. The voltage sensors are resistive dividers.

1.3	Measurand processing

Hardware

The multiplication and integration is made in a commercially available energy measuring circuit TDK71M6513.

Software

The validation of software is based on the essential requirements given in MID and WELMEC Guide 7.2. A report with number MTexxx, dated 2006-xx is issued and is held by SP.

1.4 	Indication of the measurement results

There is a LCD display for the measured electrical energy. There is also a power line communication device capable to communicate e.g. the measured energy over the power line. Only the display is part of the assessment.

1.5	Optional equipment and functions subject to MID requirements

- none

1.6	Technical documentation

-	Technical Description And Operation Manual, ADDM.411152.247 TD or ADDM.411152.245 TD

-	User Manual, ADDM.411152.247 UM or ADDM.411152.245 UM

-	Functional Description, ADDM.411152.247 FD or ADDM.411152.245 FD

1.7	Integrated equipment and functions not subject to MID

-	Power line communication device

-	230 V/ 5 A relay output

-	Optical communication port

-	Optional 80 A load switch

-	Reactive energy measurement

2	Technical data

2.1	Rated operating conditions

Measurand

Electrical energy, in kWh, with a resolution of 0,1 kWh. Three-phase four-wire direct connected applications. Phase voltages of 230 ±10 %, and currents from 0,25 A to 80 A (reference current 5 A). Power factors from 0.8 capacitive to 1 to 0.5 inductive. (Extra measurand not subjected to MID: Reactive energy)

Accuracy class

B, (Not subjected to MID: Reactive energy measurement, class 2)

Environments / influence quantities

Climatic:	-40 to +70C, closed location, condensing possible

Mechanic:	class M1

Electromagnetic:	class E2

Durability period estimated by the manufacturer

Durability period estimated to 240 months.

2.2	Other operating conditions

The LCD is switched off if the internal temperature is below -30 or exceeds +70. The optional load switch can be set to switch off the load at a number of conditions: bad voltage, bad power factor, high current, high power, bad balance and differential current.

3	Interfaces and compatibility conditions

The instrument may use the following interfaces for data communication:

Infrared Optical port, protocol proprietary of manufacturer.

Power line communication, (≤ 0,4 kV), protocol proprietary of manufacturer.

4	Requirements on production, putting into use and utilisation

4.1	Requirements on production

No special requirements identified.

4.2	Requirements on putting into use

No special requirements identified.

4.3	Requirements for consistent utilisations

No special requirements identified.

5	Control of the measuring tasks of the instrument in use

5.1	Documentation of the procedure

Test output (red LED) is provided, impulse constant is 1000 imp./kWh. Standard reference equipments can be used. Extra decimal of resolution provided by 0,01 kWh.

Voltage and current circuits can not be disconnected from each other.

5.2	Special equipment or software, if applicable

No special requirements identified.

5.3	Identification of

-	Hardware: PCB no is 758727.548-4.

-	Software: The software version is shown as the first displayed information at power-up of meter. The version approved is 5.4.08.

5.4	Calibration-/adjustment procedure, if applicable

No special requirements identified.

6	Security measures

6.1 	Sealing

Interfaces

Means for sealing of all input/output terminals are provided.

Indicator

The display is fully protected by the case sealing.

Meter case

The screws (lower) that are locking the meter case shall be sealed with wire.

Marking plate

The marking plate is placed on the meter case, it is not protected by the seal but is indelible, the serial number is on the marking plate and can be read from memory via the optical communication port.

6.2 	Data logger

Not applicable.

7	Labelling and inscriptions

7.1	Information to be borne by and to accompany the instrument

The meter shall bear the following information:

-	manufacturer’s name or trade mark and, if required, the place of manufacture;

- 	designation of type  and space for approval marks: number of the European Community (EC) type examination certificate;

-	the number of phases and the number of wires for which the meter is suitable (three-phase 4-wire)

-	the serial number and year of manufacture. If the serial number is marked on a plate fixed to the cover, the number shall also be marked on the meter base or stored in the meters non-volatile memory;

-	the reference voltage in one of the following forms: the  number of elements if more than one, and the voltage at the meter terminals of the voltage circuit(s); or, the rated voltage of the system or the secondary voltage of the instrument transformer to which the meter is intended to be connected;

-	the current measuring range:  the minimum current (Imin), reference current (Iref), and maximum current (Imax), for example 0,25-5(80) A;

-	the reference frequency in Hz;

-	the meter constant;

-	the accuracy class index of the meter;

-	the operating temperature range for which the meter is intended, or the environmental class.

-	for insulating encased meters of protective class II, the sign of the double square (IEC 60417-5172:Class II equipment.)

7.2	Conformity marking in accordance to article 17

The instrument shall be marked in accordance to article 17 which e.g. describes the CE-marking together with M, year of marking and the notified body number.

7.3	Further inscriptions, if necessary

No further inscriptions considered necessary.