T E S T     R E P O R T

Integrated Metering System

##### SMART IMS

### TESTING OF REMOTE MEASUREMENT DEVICES

#### 1  SUBJECT

The testing of remote measurement devices for compliance with technical specifications                     PT MD 17-02744093-011:2003 as referenced to IEC 61036 “Alternating current static watt-hour meters for active energy” (classes 1 and 2)”.

- DESCRIPTION OF SAMPLE

  - There were tested three single- phase meters.

The single -phase meter of electrical energy for direct connection are intended for measuring of active electrical energy in two-wire single-phase electrical alternating current networks.

The devices were marked as shown in table 1.

Table 1

| Name of device | Serial No. | Type | Rated voltage, V | Ib, A | Imax, A | Class index |
| --- | --- | --- | --- | --- | --- | --- |
| Single- phase meter for electrical energy | 00207691 | NP-10 D SM.1F | 230 +15% -20% | 5 | 65 | 1 |
|  | 00207692 |  |  |  |  |  |
|  | 00207694 |  |  |  |  |  |

##### SELECTION OF SAMPLE

The samples were selected by the customer.

##### METHOD OF TEST

  - The meters were tested for the requirements for an indoor meter in accordance with technical specifications PT MD 17-02744093-0112003 as referenced to IEC 61036.

  - All equipment used for accuracy testing is traceable to national standards.

###### TEST CONDITIONS

5.1	Temperature			(23 ± 2) ºС

5.2	Relative humidity			(60 ± 15) %

5.3	Air pressure			86 kPа.

- RESULTS

  - Constructional requirements (subsections 4.1, 4.2, 4.3)

Table 2

| Property | Result |  | Requirement of IEC 61036 |
| --- | --- | --- | --- |
| Standard electrical values: |  |  |  |
| reference voltage: 230 V | Complied | Clause 4.1 |  |
| basic current: 5 A; |  |  |  |
| maximum current: 65 A; |  |  |  |
| reference frequency: 50 Hz |  |  |  |
| Mechanical requirements: |  | Clause 4.2 |  |
| General | Complied | Clause 4.2.1 |  |
| Case | Complied | Clause 4.2.2 |  |
| Window | Complied | Clause 4.2.3 |  |
| Terminals | Complied | Clause 4.2.4 |  |
| Protective earth terminal | Not applicable |  |  |
| Terminal cover | Complied | Clause 4.2.5 |  |
| Clearance and creepage distances | Complied 1) | Clause 4.2.6 |  |
| Protective class II | Complied | Clause 4.2.7 |  |
| Resistance to heat and fire | Complied | Clause 4.2.8 |  |
| Protection against dust and water (IP 51) | Complied | Clause 4.2.9 |  |
| Solar radiation | Complied | Clause 4.2.10 |  |
| Display | Complied | Clause 4.2.11 |  |
| Output device | Complied | Clause 4.2.12 |  |
| Marking | Complied | Clause 4.2.13 |  |
| Climatic conditions | Complied | Clause 4.3 |  |
| 1) The impulse test was conducted at 12 kV |  |  |  |

  - Electrical requirements (subsection 4.4)

Table 3

| Property | Result | Requirement of IEC 1036 |
| --- | --- | --- |
| Power consumption | Complied | Clause 4.4.1 |
| Influence of supply voltage | Complied | Clause 4.4.2 |
| Short time over currents | Complied | Clause 4.4.3 |
| Influence of self heating | Complied | Clause 4.4.4 |
| Heating | Complied | Clause 4.4.5 |
| Insulation | Complied | Clause 4.4.6 |
| Electromagnetic compatibility: |  | Clause 4.5 |
| fast transient burst | Complied |  |
| electromagnetic HF fields | Complied |  |
| electrostatic discharges | Complied |  |
| radio interference | Complied |  |
| Accuracy and influence quantities |  | Clause 4.6 |
| Limits of error due to variation of current | Complied |  |
| Voltage variation | Complied |  |
| Frequency variation | Complied |  |
| DC component in the a.c. current | Complied |  |

Table 3 (continuation)

| Property | Result | Requirement of IEC 1036 |
| --- | --- | --- |
| Harmonic tests | Complied | Clause 4.6 |
| Continuous magnetic induction | Complied |  |
| Magnetic induction 0,5 mT | Complied |  |
| Temperature variation | Complied |  |
| Starting and running | Complied |  |
| Meter constant | Complied |  |

NOTE – The results obtained for the accuracy and power consumption tests are given in annex A to this report.

###### CONCLUSION

The samples complied with all requirements of the technical specifications against which it was tested.