Исходный текст начинается на 16-й странице.

Руководство по эксплуатации

SM200

Почвенный влагомер

SM200-UM-1.1

Delta-T Devices Ltd

___________________________________________________________________________

Примечания

Авторское право

Вся документация, касающаяся SM200, а также его дизайн, являются эксклюзивной собственностью компании Delta-T Devices и охраняется законом об авторском праве, © 2006.

Данный продукт находится в стадии патентования.

Соответствие требованиям Евросоюза

Продукт SM200 соответствует требованиям Евросоюза, а именно EN61326:1997.

Изменение дизайна

Компания Delta-T Devices Ltd оставляет за собой право менять дизайн, а также детали своих продуктов в любое время без предварительного уведомления.

Версия инструкции пользователя: SM200-UM-1.1 Май 2006

Delta-T, Ltd

Адрес: 128 Low Road, Барвелл

Кембридж CB5 0EJ

Соединённое королевство

Тел: +44 1638 742922

Факс: +44 1638 743155

Электронная почта: sales@delta-t.co.uk

Сайт: www.delta-t.co.uk

___________________________________________________________________________

Оглавление

Вступление 4

Описание 4

Комплектация 5

Меры предосторожности 6

Как работает SM200 7

Обращение 8

Кабели 8

Установка 9

Схема подключения датчика 10

GP1 10

DL6 11

DL2e 12

Другие датчики данных 13

Измеритель влажности HH2 14

Калибровка 15

Калибровка почвы 16

Калибровка сенсора 18

Считывание данных 19

Выявление и устранение неполадок 20

Технические характеристики 21

Спецификации 21

Определения 25

Ссылки 27

Техническая поддержка  28

Калибровка по типу почвы 30

Лабораторная калибровка для неглинистых почв 31

Лабораторная калибровка для глинистых почв 33

Предметный указатель 36

Стр 3

Вступление

Описание

Влагомер SM200 измеряет влажность почвы в заданном месте на основе анализа приблизительно 0.5 литра почвы (либо другого субстрата). Он покрыт герметизированным пластиковым корпусом, имеет диаметр ~40мм, и два выведенных измерительных зонда. Датчик также имеет влагонепроницаемый коннектор, соединённый с выводными кабелями.

Анализ почвы производится, когда штыри воткнуты непосредственно в почву.

SM200 работает при постоянном напряжении. Данный продукт может быть применён в целях измерения влажности почвы определённого типа при помощи общих калибровок, приведенных ниже, либо калибровок для какого-либо определённого типа почвы.

Характеристики

􀂄 Погрешность измерений ± 3%

􀂄 Низкая чувствительность к засоленности

􀂄 Отличная устойчивость по отношению к температуре

􀂄 Минимальное влияние на почву

􀂄 Простая установка на глубину в заранее заготовленные отверстия

􀂄 Коннектор IP68

􀂄 Прочный, влагозащищённый, помещаемый в землю

Стр 4

Комплектация

Ваш комплект может быть укомплектован следующим образом:

| Комплектующие | Код | Описание |
| --- | --- | --- |
|  | SM200 | Влагомер SM200, оснащённый коннектором M8, в защитном корпусе |
|  | - (оснащённый SM200) | Проводник быстрого запуска |
|  | SM2C/d- HH2 | Кабель (1,1 м) для подключения SM200 к HH2 при помощи 25-штыревого D-конвектора |
|  | SM2C/w-05 | 5-метровый кабель для отключения SM200 к регистраторам. На конце имеет оголённые провода |
|  | EXT/M8-05 EXT/M8-10 EXT/M8-25 | Удлинители длиной 5, 10 и 25 метров. На концах имеют inline-коннекторы IP68 |
|  | SM-INRD1 | Вводной зонд длиной 1 м |
|  | ML2-AG1- SL | Спиральный бур диаметром 45мм и длиной 1,2м |

Стр. 5

Меры предосторожности

| 􀂄 Зонды, входящие в комплект, SM200, являются острыми, что упрощает их введение. В целях безопасности следует придерживаться следующих рекомендаций | В целях личной безопасности, а также во избежание повреждения зондов следует хранить и транспортировать SM200 в защитном корпусе Внимание! острые зонды Придерживайтесь мер предосторожности |
| --- | --- |
| 􀂄 Избегайте прикосновения к зондам, а также иных действий, которые могут стать причиной их повреждения. Держите SM200 в защитном корпусе, когда не используете его. |  |

􀂄 Перед подсоединением кабелей убедитесь, что коннекторы являются чистыми, не повреждёнными, а также в том, что они правильно направлены.

􀂄 Прикручивайте коннекторы крепко во избежание попадания влаги. В случае если коннектор повреждён, смажьте его слегка силиконовой смазкой в целях уплотнения соединения.

􀂄 Не вытягивайте SM200 из почвы за кабель.

􀂄 Если вы чувствуете сильное противодействие при введении SM200 в почву, проверьте, нет ли там камня. В таком случае следует выбрать другое место введения.

Стр. 6

Как работает SM200

|  | Когда SM200 подключён к сети,.. |
| --- | --- |
|  | ...частота волн составляет 100MГц (как в FM- диапазоне). |
|  | Ток такой же частоты проходит по стальным зондам, по которым электромагнитное поле распространяется в почву. |
|  | Содержание воды в почве, в которую помещён зонд,... |
|  | ...оказывает влияние на её диэлектрическую проницаемость. (меру реакции материи на поляризацию в электромагнитном поле). Диэлектрическая проницаемость воды составляет ≈ 81, Тогда как диэлектрическая проницаемость почвы ≈ 4, а воздуха - ≈ 1 |
|  | Диэлектрическая проницаемость почвы имеет сильное влияние на поле… |
| Ввых | …регистрируемое SM200, обуславливая стабильное выходное напряжение… |
| Влажность почвы 22 % | ...являющееся простым, точным показателем содержания влаги в почве. |

Стр. 7

Применение

Кабели

| Кабель HH2 | Удлинители | Кабель регистратора |
| --- | --- | --- |
| 1.1м | 5м 10м 25м | 5м |

􀂄 Если вы собираетесь установить SM200 на глубину, используя вводной штырь, проденьте

Коннектор сквозь штырь до подсоединения кабеля.

􀂄 Придерживайтесь мер предосторожности во время подсоединения кабелей. Убедитесь, что коннекторы чистые, не имеют повреждений, а также что они направлены должным образом перед тем, как производить подсоединение.

􀂄 Соедините их крепко, чтобы избежать попадания влаги.

􀂄 Суммарная длина удлинителей не должна превышать 55 метров.

Предназначение кабелей

SM200 – это одноконечный single-ended сенсор.

Кабели имеют следующее назначение:

| Функции | Название | Цвет |
| --- | --- | --- |
| Напряжение, подаваемое на сенсор | напряжение U+ | Красный |
| Выход сенсора + | Сигнал | Белый |
| Напряжение U- Выход сенсора – | Заземление | Защитный (серебренный) |

Стр. 8

Установка

Установка на поверхности и подготовка установочной площадки

􀂄 Уберите все камни. Если почва является очень твёрдой, приготовьте заранее отверстия для введения зондов.

􀂄 Вводите SM200 в почву до тех пор, пока штыри не окажутся полностью погружёнными в неё. Обеспечьте хороший контакт с почвой.

􀂄 Если вы чувствуете сильное сопротивление во время введения штыря в почву, то, вероятно, там находится камень. В таком случае следует выбрать другое место.

Установка на глубину

􀂄 Пробуравьте отверстие диаметром 45мм под углом ~10° к нормали.

􀂄 Подсоедините вводной штырь к SM200. Помните, что прежде нужно продеть кабель сквозь вводной штырь и подсоединить коннектор.

􀂄 Вводите SM200 в почву, пока штыри не окажутся полностью погружёнными в неё. Обеспечьте хороший контакт с почвой.

􀂄 Заполните грунтом пространство вокруг штыря.

Альтернативный вариант

􀂄 Выройте канаву и произведите горизонтальную установку (как показано на рисунке).

Стр. 9

Подключение и конфигурация регистратора

GP1

Два SM200 могут быть подсоединены к каждому GP1, как показано на рисунке.

Рисунок и табличка демонстрируют подключения к Каналу 1:

| Проводка SM200 | Цвет | Терминал GP1 |
| --- | --- | --- |
| Напряжеине U+ | Красный | CH1 (PWR) |
| Сигнал | Белый | CH1 (+) |
| Заземление | Серебренный | CH1 (GND) и CH1 (-) |

Переключите GP1, выбирая “SM200” в меню Delta LINK-PC или Pocket Delta LINK.

Подробнее об этом читайте в проводнике (GP1 Quick Start Guide).

Стр. 10

DL6

К каждому DL6 может быть подключено до шести SM200.

Рисунок и табличка демонстрируют подключения к Каналу 1:

| Проводка SM200 | Цвет | Терминал DL6 |
| --- | --- | --- |
| Напряжение U+ | Красный | U+ |
| Сигнал | Белый | IN+ |
| Заземление | Серебренный | IN- и (0V) |

Переключите DL6, выбирая “SM200” в меню Delta LINK-PC или Pocket Delta LINK. Подробнее об этом читайте в проводнике (DL6 Quick Start Guide).

Стр. 11

DL2e

Переключите DL2e, выбирая один из кодов сенсора SM200 в библиотеке сенсоров LS2Win, в соответствии с вашим типом почвы.

К каждому DL2e может быть подключено до шестидесяти SM200.

Рисунок и табличка демонстрируют подключения к Каналу 58 с использованием платы обработки входящих сигналов LAC1, настроенную на канал 15, и предварительный канал 63:

| Проводка SM200 | Цвет | Терминал DL2e |
| --- | --- | --- |
| Напряжеине U+ | Красный | 63 NO |
| Сигнал | Белый | 58+ |
| Заземление | Серебренный | 58- и 61- |

Подробнее об этом читайте в пользовательской документации по DL2e.

Стр. 12

Другие регистраторы данных

􀂄 SM200 следует подключить как одноконечный сенсор, подключённый к сети.

􀂄 Настройте вход (input) регистратора как сенсор напряжения, приблизительно в пределах от 0 до 1,0 Вольт, что соответствует содержанию воды приблизительно от 0 to 60% (см. след. часть).

􀂄 Оптимальный период выдержки для SM200 составляет приблизительно в 1 секунду. Не рекомендуется постоянно держать сенсор под напряжением.

􀂄 Настройте регистратор на преобразование данных SM200 во влажность почвы, используя нелинейные либо линейные графики (см. Калибровку).

Стр. 13

Измеритель влажности HH2

􀂄 Подключите SM200 к измерителю HH2.

􀂄 Нажмите Esc, чтобы включить измеритель. В случае необходимости повторите это действие, пока на экране HH2 не отобразится начало работы.

􀂄 Настройте измеритель на считывание информации с SM200:

► Нажмите Set и перейдите по списку ниже до опции Device (Устройство).

► Снова нажмите Set и дойдите до SM200:

Device:

SM200

► Нажмите Set, чтобы подтвердить выбор.

􀂄 Если вы хотите сохранить данные, то можете определить каждое данное при помощи схематических пометок и нумерации. Для этого нужно нажать Set и перейти к опции Data.

􀂄 Нажмите Read для считывания данных (это займёт около 2 секунд).

Вы можете выбрать разные разделы в опции Display.

SM200 Store?

20.3%vol

􀂄 Нажмите Store для сохранения либо Esc для сброса данных.

􀂄 Переместите SM200 в другое место...

Если вы сохранили данные, подключите ваш HH2 к PC и запустите

HH2Read, чтобы извлечь данные.

Подробнее об этом читайте в инструкции (HH2 User Manual).

Стр. 14

Калибровка

SM200 имеет общие калибровки для минеральных и органических почв. Эти калибровки могут быть использованы для непосредственного преобразования выходного сигнала с сенсора во влажность почвы при помощи использования регистраторов Delta-T и измерителя влажности HH2. Данный раздел поясняет, как работают эти калибровки и как использовать их применительно к другим типам почвы, а также как использовать калибровки с другими регистраторами.

SM200 измеряет объёмную влажность почвы θ посредством определения диэлектрических свойств влажной почвы: диэлектрическую проницаемость ε либо более подходящую величину, коэффициент рефракции, который почти эквивалентен √ε. Последовательность измерений SM200 более наглядно изображена ниже:

| 1. Калибровка почвы | + | 2. Калибровка сенсора |
| --- | --- | --- |
| θ → √ε |  | U → √ε |
| Калибровки почвы График √ε (θ) |  | Диэлектрические характеристики SM200 График √ε (U) |

3. Показания влажности почвы

U → θ

Показания влажности почвы SM200

График θ (U)

Стр. 15

Калибровка почвы

Влажная почва представляет собой в основном смесь частиц грунта, воздуха и воды. В совокупности все эти составляющие определяют её диэлектрические свойства, включая коэффициент рефракции  √ε. При вычислении коэффициента рефракции смеси учитываются вклады отдельных компонентов [ссылка 4].

Для определённого типа почвы такой вклад может приниматься за константу. Таким образом, изменение коэффициента рефракции, измеренного при помощи SM200, может быть вызвано только изменением влажности почвы θ. Данное соотношение можно выразить следующим образом:

√ε = a0+a1* θ, где коэффициенты a0 и a1 являются параметрами диэлектрических свойств разных типов почвы.

График Калибровки почвы

Ось X – Влажность почвы (м3.м-3)

Ось Y – Коэффициент рефракции почвы (√ε)

Смещение (a0)

Угловой коэффициент (a1)

Следует также заметить, что:

a0 = √εсух_почва обычно находится в пределах от 1,3 до 2,3

a1 приблизительно соответствует √εводы – 1 и обычно составляет приблизительно 8,0.

Реальные значения a0 и a1 могут значительно отличаться от приведённых выше в том случае, если они обусловлены дополнительными факторами. В частности, связанная вода, содержащаяся в глинистых грунтах, может обуславливать более высокие значения a1.

Стр. 16

Общие калибровки почвы

Большинство типов почв может быть охарактеризовано при помощи одной или двух калибровок, предлагаемых ниже. Одной – для минеральных грунтов (в основном, для песка, илистых грунтов и глины) и одной – для органических почв (с высоким содержанием органических веществ).

|  | a0 | a1 |
| --- | --- | --- |
| Минеральный грунт | 1,6 | 8,4 |
| Органический грунт | 1,3 | 7,7 |

График  Общие калибровки почвы

Ось X - Влажность почвы (м3*м-3)

Ось Y - Коэффициент рефракции почвы (√ε)

Красный – минеральная

Зелёный - органическая

Данные значения были использованы для обобщения нелинейных преобразований и линеаризационных таблиц в разделе Получение значений влажности почвы.

Специальные калибровки почвы

Вместо применения общих калибровок вы можете определить калибровочные величины a0 и a1 конкретно для вашего типа почвы. Эта процедура довольно проста в случае, если у вас есть доступ к стандартному лабораторному оборудованию, и детально описана в приложении А.

Специальная калибровка почвы может значительно повысить точность считывания конкретных данных, но малоэффективна в случае, когда допущены значительные ошибки при подборе образцов и установке.

Стр. 17

Калибровка сенсора

Каждый SM200 приспособлен специально для получения конкретных диэлектрических характеристик:

График  Диэлектрические свойства SM200

Ось X – Выход SM200 (В)

Ось Y - Коэффициент рефракции почвы (√ε)

Данный график может быть аппроксимирован либо при помощи многочлена, либо при помощи линеаризационной таблицы:

Многочлен (используется для данных SM200, лежащих в промежутке от 0 до 1,1 В)

√ε = 1,0+16,103U – 38,725U2 + 60,881U3– 46,032U4+ 13,536U5

Где В - выход SM200 в Вольтах

Линеаризационная таблица  (используется для данных SM200 от 0 В до 1,2 В)

| В | √ε | В | √ε | В | √ε | В | √ε |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0,000 | 1,011 | 0,150 | 2,729 | 0,400 | 4,102 | 1,100 | 7,290 |
| 0,030 | 1,453 | 0,180 | 2,956 | 0,600 | 4,956 | 1,150 | 7,610 |
| 0,060 | 1,842 | 0,220 | 3,218 | 0,820 | 5,939 | 1,200 | 7,982 |
| 0,090 | 2,179 | 0,260 | 3,446 | 0,960 | 6,571 | 1,220 | 8,205 |
| 0,120 | 2,473 | 0,310 | 3,701 | 1,040 | 6,960 | 1,800 | 9,000 |

Стр. 18

Получение значений влажности почвы

Нелинейное преобразование

Используя методы калибровки почвы и сенсора, получаем следующее уравнение:

θ=((1,0+16,103U – 38,725U2 + 60,881U3– 46,032U4+ 13,536U5)- a0)/ a1

где a0 и a1 – калибровочные коэффициенты.

Для среднестатистической минеральной почвы получаем:

θминер= - 0,071+1,917U – 4,6U2 + 7,248U3– 5,48U4+ 1,611 U5

А для среднестатистической органической почвы получаем:

θминер= - 0,039 + 2,091U – 5,029U2 + 7,907U3– 5,978U4+ 1,758U5

Преобразование посредством линеаризационной таблицы

Преобразование данных SM200 (В) во влажность почвы θ (м3*м-3 или объёмный процент) может быть также произведено с помощью линеаризационной таблицы.

Данная таблица отображает значения, используемые для регистратора данных DL2e:

| Влажность почвы | Минеральная почва | Органическая почва | Влажность почвы | Минеральная почва | Органическая почва |
| --- | --- | --- | --- | --- | --- |
| Объёмные проценты | В | В | Объёмные проценты | Volts | В |
| -4 | -2,090 | -2,090 | 52 | 0,826 | 0.678 |
| 0 | 0,041 | 0,019 | 56 | 0,901 | 0,747 |
| 4 | 0,068 | 0,041 | 60 | 0,975 | 0,816 |
| 8 | 0,099 | 0,066 | 64 | 1,043 | 0,884 |
| 12 | 0,135 | 0,094 | 68 | 1,103 | 0,952 |
| 16 | 0,178 | 0,126 | 72 | 1,156 | 1,017 |
| 20 | 0,230 | 0,164 | 76 | 1,199 | 1,076 |
| 24 | 0,292 | 0,209 | 80 | 1,237 | 1,127 |
| 28 | 0,365 | 0,262 | 84 | 1,269 | 1,171 |
| 32 | 0,443 | 0,324 | 88 | 1,297 | 1,210 |
| 36 | 0,523 | 0,393 | 92 | 1,322 | 1,244 |
| 40 | 0,601 | 0,466 | 96 | 1,347 | 1,278 |
| 44 | 0,676 | 0,538 | 100 | 1,800 | 1,800 |
| 48 | 0,751 | 0,609 | 104 | 2,090 | 2,090 |

Стр. 19

Выявление и устранение неисправностей

Всегда пытайтесь выяснить, какая именно часть измерительной установки является причиной неполадки. В случае с SM200 возможны следующие варианты:

Измерительное устройство

Какое оборудование используется для считывания данных?

• Измеритель влажности Delta-T HH2

• Один из регистраторов данных Delta-T: GP1, DL6 or DL2e

Обратитесь к инструкции по пользованию данными устройствами, либо поищите формацию и программное обеспечение, относящееся к использованию данных устройств, в сети Интернет. Попытайтесь воспользоваться альтернативным оборудованием, если имеете доступ к таковому.

Убедитесь в том, что используете соответствующие калибровки почвы и подходящий метод преобразования (см. раздел «Калибровка»).

SM200

Попытайтесь определить, не являются ли причиной неполадок:

• Непосредственно SM200 либо соединительный кабель

Либо:

• Механические проблемы либо повреждения

• Электрические неполадки либо неполадки на электронном уровне

Проверка функционирования системы

Для проверки правильности функционирования SM200 могут быть использованы два следующих простых метода

Считывание данных из воздуха

Держите SM200 на удалённом расстоянии от других предметов и попытайтесь снять данные, используя измеритель HH2, либо другой измеритель или регистратор. Данные должны находиться в пределах 0±12мВ при условии использовании пятиметрового кабеля

Считывание данных с пальцев

Обхватите пальцами SM200 крепко (но аккуратно!) вокруг двух металлических зондов и снимите данные, используя измеритель HH2, либо другой измеритель или регистратор. Данные должны находиться в пределах от 500 до 750 мВ.

Стр. 20

Технические характеристики

Спецификации

|  |  | Диапазон и примечания |  |
| --- | --- | --- | --- |
| Погрешность* | ± 3 %vol | от 0 до 50 %vol |  |
| Погрешности минерализации | ± 3,5 %vol | от 0 до 40 %vol, и выше | от 50 до 500 мS*м-1 (от 0,5 до5 дS*м -1) |
| Необходимые характеристики сети | от 5 до 14В 15мA | Рекомендуемое время выдержки – 1 секунда (мин – 200мс) |  |
| Выходной сигнал | от 0 до 1,0В | Номинально от 0 до 60 %vol |  |
| Уплотнение среды | IP68 | Сенсор, помещаемый в грунт, герметично соединённый с IP68 |  |
| Условия работы | от -20 до +60°C | Данные, считанные с промёрзшей почвы, не являются действительными |  |

*Данные погрешности применимы к калибруемым почвам в условиях от 0 to 60°C.

Данные, указанные в процентах,(%vol),, являются объёмным содержанием воды.

Электропроводность

График SM200 Электропроводность при различном содержании влаги

Ось X Электропроводность ECp (мS/м)

Ось Y Выход SM200 (В)

100%(вода)

Розовая кривая: В условиях типичного искусственного грунта

Зелёная кривая: В условиях обычного грунта

Стр. 21

Температурная зависимость

SM200 температурная зависимость для каждого определённого типа почвы будет определяться рядом факторов:

􀂄 Электроника SM200 имеет очень низкую температурную зависимость и вносит пренебрежительно малый вклад в общую чувствительность.

􀂄 Коэффициент рефракции воды (√ε (см. Раздел «Калибровка»)) понижается с ростом температуры. Это обуславливает обратную температурную зависимость, особенно в почвах либо субстратах с высоким содержанием влаги.

􀂄 Вода, связанная с поверхностью частиц почвы, имеет гораздо более низкий коэффициент рефракции, нежели свободная вода. Процент связанной воды повышается с ростом температуры, что обуславливает прямую температурную зависимость, особенно в почвах с низким содержанием влаги.

Два последних эффекта частично компенсируют друг друга, однако в почве, где доминирует один из этих эффектов, SM200 будет иметь значительную температурную зависимость. Данная иллюстрация основана на модели, показанной в ссылке 7:

График Иллюстрация температурной зависимости данных SM200 в глинистой почве

Ось X Температура почвы (°C)

Ось Y Содержание влаги в почве (м 3*м-3)

Розовая линия – связанная вода

Note: коэффициент рефракции льда значительно отличается от коэффициента рефракции воды, поэтому данные SM200

не будут являться достоверными в случае, если измерения проводятся в почве, где температура опустилась ниже 0°C.

Стр. 22

Тестируемый объём

SM200 наиболее чувствителен к сигналам в непосредственной близости к зондам, однако небольшая доля сигнала распространяется на расстояние до 50мм от них. Чтобы достичь наибольшей точности, рекомендуется брать данные с объёма почвы приблизительно 0,5 – 1 л.

SM200 могут оказывать влияние на работу друг друга в случае, если они расположены слишком близко один к другому. Поэтому их следует размещать по крайней мере на расстоянии 100мм друг от друга.

Область чувствительности для SM200, помещённого в воду, отображена на диаграмме (вертикальная ось имеет произвольную размерность):

График Область чувствительности SM200 вблизи зондов

Ось X Расстояние от центральной линии между зондами в направлении, параллельном плоскости зондов (мм)

Ось Z Расстояние от плоскости зондов в направлении, перпендикулярном этой плоскости (мм)

Стр. 23

Электромагнитная совместимость (EMC)

Общая информация

SM200 – Продукт класса А, предназначенный для использования вне страны производства.

Используйте кабели и аксессуары, рекомендованные компанией Delta-T (другие кабели могут неблагоприятно повлиять на работу устройства и стать причиной получения недостоверных данных).

Если возможно, проведите кабели вдоль поверхности, либо спрячьте их в грунт (это уменьшит вероятность их повреждения людьми или животными). Не вносите никаких изменений в модификацию продукта либо прилагаемых к нему аксессуаров.

Нормативная информация

Европа

SM200 был одобрен CE в рамках директивы 89/336/EEC Европейского Союза EMC для использования согласно данной инструкции. Использование кабелей либо аксессуаров, не рекомендованных компанией Delta-T, выходит за рамки данной директивы.

Соответствие FCC (США)

Данное устройство удовлетворяет 15-ти правилам FCC. Его использование должно отвечать определённым условиям:

1. Устройство не должно причинять вреда

2. Должна допускаться возможность вмешательства, включая нежелательное.

Следует отметить, что данное оборудование было протестировано и отмечено как соответствующее уровню приборов класса А., в соответствии с частью 15 правил FCC.

Данные правила были разработаны в целях обеспечения определённой защиты от вредоносного вмешательства при использовании оборудования в коммерческих целях. Данное оборудование является энегропотребляющим и энергопроизводящим, а также может излучать энергию с частотой радиоволн в случае, если не следовать инструкции, что может быть причиной вредного влияния на радиокоммунакиции. Использование данного оборудования в жилых районах может оказывать вредное воздействие, которое пользователь должен будет устранять при помощи собственных средств.

Использование модификаций, не одобренных изготовителем, может не соответствовать правилам FCC.

Стр. 24

Определения

Объёмное содержание влаги в почве  определяется как

Θu= Vвлажн/Vобр

где Vвлажн - объем воды, содержащейся в образце

а Vобр  - полный объём образца почвы.

Предпочитаемой единицей измерения в данном случае является [м3 * м-3], однако [%vol] также часто используется. Содержание влаги в почве обычно находится в пределах от 0,02 м3*м-3 и ниже для песчаных почв, около 0.4 м3*м-3 для глинистых почв и достигает во влажных торфяных почвах 0.85 м3*м-3.

Гравиметрическое содержание влаги в почве определяется как

Θg= (mвлажн/mобр)*g*g-1

Где  mвлажн  - масса воды в образце,

а mобр  - полная масса сухого образца.

Для перехода от объёмного содержания влаги в почве к гравиметрическому используется следующее соотношение

Θg= Θu *(ρвлажн/ρобр)

где ρвлажн  - плотность воды (= 1г* cm-3),

а ρобр – удельная плотность образца (mобр/Vобр)

Показатели органической и минеральной почвы:

При помощи общих калибровок были выведены следующие показатели для широкого спектра типов почв:

| Тип почвы | оптим. содерж. органики: | исп. для содерж. органики: | диапазон удельной плотности: (г.см-3) | исп. для удельн. плотности: (г.см-3) |
| --- | --- | --- | --- | --- |
| Минеральная | ~ 1 %C | <7 %C | 1,25 – 1,5 | > 1,0 |
| Органическая | ~ 40 %C | > 7 %C | 0,2 – 0,7 | <1,0 |

Стр. 25

Соленость почвы

Предпочтительные единицы измерения СИ для ионной проводимости – [мS*м-1] (где S - это

Siemens, единица электропроводности, измеряемая также в [Ом-1]).

В данном случае применимы следующие соотношения:

1 мS*м-1 = 0,01 дS*м-1

= 0.01 мS*см-1

= 10 µS*см-1

Типы почв можно разделить по показателю солёности следующим образом:

Несолёные                           0 - 200 мS*см-1

Незначительно солёные    200 - 400 мS*м-1

Умеренно солёные             400 - 800 мS*м-1

Значительно солёные      800 - 1600 мS.м-1

Чрезвычайно солёные  > 1600 мS*м-1

Стр. 26

Литература (ссылки)

1. Gaskin, G.J. and J.D. Miller, 1996

Measurement of soil water content using a simplified impedance

measuring technique.

J. Agr. Engng Res 63, 153-160

2. Topp, G.C., J. L. Davis and A. P Annan 1980

Electromagnetic determination of soil water content.

Water Resour. Res 16(3) 574-582

3. Whalley, W.R. 1993

Considerations on the use of time-domain reflectometry (TDR) for

measuring soil moisture content.

Journal of Soil Sci. 44, 1-9

4. White, I., J.H. Knight, S.J. Zegelin, and Topp, G.C. 1994

Comments on ‘Considerations on the use of time-domain

reflectometry (TDR) for measuring soil water content’ by W R

Whalley

Journal of Soil Sci. 45, 503-508

5. Roth, C.H., M.A. Malicki, and R. Plagge, 1992

Empirical evaluation of the relationship between soil dielectric

constant and volumetric water content as the basis for calibrating

soil moisture measurements.

Journal of Soil Sci. 43, 1-13

6. Knight, J.H. 1992

Sensitivity of Time Domain Reflectometry measurements to lateral

variations in soil water content.

Water Resour. Res., 28, 2345-2352

7. Or, D. and J.M. Wraith 1999

Temperature effects on soil bulk dielectric permittivity measured by

time domain reflectrometry: A physical model.

Water Resour Res., 35, 371-383

Стр. 27

Техническая поддержка

Условия Продажи

Наши условия продажи (см.: Усл: 1/00) определяют юридические обязательства компании Delta-T относительно данного продукта. Следующие пункты обобщают позицию компании Delta-T , однако в каждом конкретном случае необходимо ссылаться на конкретные условия продажи, которые будут иметь преимущество над следующими положениями.

Компания Delta-T гарантирует отсутствие дефектов, являющихся причиной использования недоброкачественных материалов либо самого качества изготовления оборудования в течение двенадцати месяцев от даты поставки.

Компания Delta-T  не несёт никакой ответственности за повреждения, вызванные неаккуратным пользованием, неквалифицированным сервисным обслуживанием, либо за повреждения, являющиеся следствием износа оборудования, а также за повреждения, вызванные причинами, не находящимися в области контроля компании Delta-T.

В случае возникновения каких-либо проблем с использованием данного оборудования пользователь должен оповестить об этом компанию Delta-T  либо её местного дистрибьютора сразу же после обнаружения данной проблемы.

Компания Delta-T может решить этот вопрос посредством бесплатного предоставления запасных частей либо посредством ремонта оборудования на территории Соединённого Королевства в течение гарантийного периода.

Однако в таком случае компания Delta-T не будет нести никаких транспортных и таможенных расходов.

Данный вопрос предпочтительно обсудить заранее, и вы можете отказаться от ваших требований на ваше усмотрение.

Компания Delta-T  не производит бесплатную поставку оборудования либо ремонт устаревшего либо снятого с производства оборудования. При этом компания Delta-T может оказать содействие в решении проблемы.

Компания Delta-T не несёт ответственности перед покупателем за косвенные убытки, повреждения товара либо в случае предоставление компенсации (независимо от того, является ли причиной небрежность компании Delta-T, её служащих, дистрибьюторов или кого-либо ещё) в результате поставки товаров и/или услуг, их использования или перепродажи покупателем.

Компания Delta-T также не несёт ответственности перед покупателем за задержку либо отказ от выполнения своих обязательств касательно своих товаров и/или услуг, если такие задержка или отказ происходили по причине, не подвластной  контролю компании Delta-T .

Стр. 28

Обслуживание и Запчасти

Пользователи, находящиеся в странах, где есть дистрибьюторы либо представители технической поддержки компании Delta-T, должны в первую очередь связываться с этими представителями по вопросам всех возникающих проблем.

Запасные части для нашего оборудования поставляются с наших заводов. Это обычно занимает несколько рабочих дней с момента получения заказа.

Поставка запасных частей и аксессуаров для датчиков или других продуктов, не производимых компанией Delta-T, занимает больше времени.

Никакие товары и оборудование не должны быть возвращены компании Delta-T без предварительного согласования с ней  либо с её дистрибьютором.

После получения оборудования компания Delta-T производит его осмотр, и сообщает пользователю о возможной стоимости ремонта и о времени, которое он займёт. Обычно это занимает несколько рабочих дней с момента получения оборудования. Однако, в случае если для осуществления ремонта либо калибровки оборудование

необходимо доставить нашему поставщику, это может занять несколько недель.

Техническая поддержка

Компания Delta-T обеспечивает техническую поддержку своим продуктам.

Пользователи, находящиеся в странах, где есть дистрибьюторы либо представители технической поддержки компании Delta-T, должны в первую очередь обратиться за такой поддержкой к ним.

Вопросы, касающиеся технической поддержки, которые получит компания Delta-T , будут обработаны её командой Технической поддержки. Вашему запросу будет присвоен определённый Т-номер, а также будет определено время, необходимое для диагностики. Т-номер вам понадобится для удобства в отслеживании процесса устранения неполадок.

При отправке запроса следует указывать серийные номера приборов, версию программного обеспечения

и приблизительную дату покупки и происхождение продукта, о котором идёт речь в запросе.

Контактная информация:

Команда технической поддержки компании

Delta-T Devices Ltd

128 Low Road, Барвелл, Кембридж, CB5 0EJ, Соединённое Королевство.

Электронная почта: tech.support@delta-t.co.uk

Веб-сайт: www.delta-t.co.uk

Тел.: +44 (0) 1638 742922

Факс: +44 (0) 1638 743155

Стр. 29

Специфическая калибровка почвы

Данное примечание детально описывает два метода получения калибровок для конкретных типов почвы.

1. Лабораторная калибровка для субстратов* и неглинистых почв

2. Лабораторная калибровка для глинистых почв

* Термин «субстрат», относится ко всем искусственным грунтам.

Основной  принцип

Содержание влаги в почве (θ) пропорционально коэффициенту её рефракции (√ε), что заложено в принцип работы SM200 (см. раздел Калибровки).

График Зависимость диэлектрических свойств почвы от содержания влаги

Ось X: содержание влаги (м3* м -3)

Ось Y: √ε

Цель калибровки состоит в нахождении коэффициентов (a0, a1), которые могут быть подставлены в линейное уравнение, с помощью которого можно получить величину влажности почвы из считываемых данных:

√ε = a0 + a1 * θ

Стр. 30

Лабораторная калибровка для неглинистых почв

Данная техника является самой лёгкой, однако она не подходит для почв, которые легко ссыхаются, в результате чего становятся очень твёрдыми.

Необходимое оборудование:

􀂄 SM200 и измеритель

􀂄 Бур для почвы (в случае достаточно вязкой почвы по сравнению с песчаной почвой или субстратом)

􀂄 Огнеупорная ёмкость (≥ 0.5 литра)

􀂄 Весы (с погрешностью <1 г)

􀂄 Печь с регулируемым температурным режимом (для минеральных почв или субстратов)

| Процесс | Примеры и замечания |
| --- | --- |
|  | Возьмите образец влажной почвы либо субстрата. Этот образец должен обладать своей естественной плотностью, иметь объём ≥ 0.5 литра и иметь правильную форму, чтобы поместиться в ёмкость. Влага в нём должна быть распределена равномерно. В случае если почва вязкая, лучше воспользоваться буром. Песчаные почвы можно легко насыпать в ёмкость, однако в таком случае необходимо незамедлительно приступать к измерениям, так как вода быстро достигнет дна ёмкости. Сжимаемые почвы и компосты часто требуют измерения своей естественной плотности в целях последующего воспроизведения её внутри ёмкости. |
|  | Измерьте объем, занимаемый образцом. Vобр = 463,5 мл |
|  | Взвесьте образец вместе с ёмкостью mвлажн = 743,3 г |

Стр. 31

|  | Вставьте SM200 в образец и запишите его показания в В. Uвлажн = 0,350В |
| --- | --- |
|  | Высушите образец полностью. В случае если почва минеральная, то рекомендуется выдержать образец в печи несколько часов или дней (в зависимости от размера и пористости образца) при температуре 105°C. В случае если почва органическая, то можно подвергнуть образец обдуванию воздухом, чтобы избежать сжигания неустойчивых фракций. |
|  | Взвесьте сухой образец в ёмкости. м0= 627,2 г |
|  | Повторно вставьте SM200 в сухой образец и запишите его показания. U0 = 0,051В |
| Вычисление a0 | Для SM200, √ε = 1,0+16,103U – 38,725U 2 + 60,881U 3 – 46,032U4 + 13,536U5 В сухой почве U = U0 = 0,051В, и при подстановке данного значение в вышеприведённое уравнение, получаем √ε 0 = 1,73 Таким образом, θ 0 = 0, а a0 =1,73 a0 = 1,73 |
| Вычисление θвлажн | Содержание воды во влажной почве, θвлажн, может быть вычислено исходя из веса воды, потерянной при высыхании, (mвлажн – m0) и её объёма, Vобр: θu= (mвлажн – m0)/Vобр = (743,3 – 627,2)/463,5 = 0,25 θu = 0,25 |

Стр. 32

| Вычисление a1 | Во влажной почве U = Uвлажн = 0,350В, а после подстановки получаем √εвлажн = 3,88 Таким образом a1 = (√εвлажн - √ε 0) / (θвлажн - θ0) = (3,88 – 1,73) / (0,25 - 0) = 8,60 a1 = 8,60 |
| --- | --- |

Лабораторная калибровка для глинистых почв

Данную технику можно использовать в том случае, когда SM200 почти невозможно ввести в сухую глинистую почву. В таком случае потребуется производить измерения при двух значительно различающихся, но ощутимых уровнях влажности.

Необходимое оборудование:

􀂄 SM200 и измеритель

􀂄 Бур для почвы

􀂄 Огнеупорная ёмкость (≥ 500 мл)

􀂄 Весы (погрешность <1г)

􀂄 Печь с регулируемым температурным режимом

| Процесс | Примеры и замечания |
| --- | --- |
|  | Возьмите влажный образец глинистой почвы (лучше, если содержание воды будет находиться в пределах от 25 до 30 %). Этот образец должен обладать своей естественной плотностью, иметь объём ≥ 500 мл, а также правильную форму, чтобы поместиться в ёмкость. Рекомендуется воспользоваться буром. Влага в нём должна быть распределена равномерно. |
|  | Измерьте объем, занимаемый образцом. Vобр = 463,5 мл |
|  | Взвесьте влажный образец вместе с ёмкостью. mвлажн = 743,3 г |

Стр. 33

|  | Вставьте SM200 во влажный образец и запишите его показания в В. Uвлажн = 0,349В |
| --- | --- |
|  | Аккуратно подсушите образец до такого состояния, чтобы содержание влаги составляло ~15 %. Следите за тем, чтобы не происходило локального пересыхания почвы. Перед снятием показаний и подержите образец ещё какое-то время, чтобы влага смогла распределиться равномерно. |
|  | Повторно взвесьте. mм = 693,2 г |
|  | Повтороно измерьте с SM200 Uм = 0,180В |
|  | Высушите образец полностью. В случае если почва минеральная, то рекомендуется выдержать образец в печи несколько часов или дней (в зависимости от размера и пористости образца) при температуре 105°C. |
|  | Взвесьте сухой образец в ёмкости. m0 = 627,2 г |

Стр. 34

|  | Подстановка в уравнение для SM200 √ε = 1,0+16,103В – 38,725В2 + 60,881В3 – 46,032В4 + 13,536В5 даёт две величины, √ε влажн и √ε м., при известных θвлажн и θм |
| --- | --- |
|  | Подстановка Vвлажн = 0.349 дает √ε влажн = 3,88 = a0 + a1 * θвлажн При θвлажн = (743,3 – 627,2)/463,5 = 0,25 |
|  | Подстановка Vм = 0,180 дает √ε м. = 2,95 = a0 + a1 * θм При θм = (693,2 – 627,2)/463,5 = 0,14 |
|  | Тогда a1 = (√ε влажн - √ε м) / (θвлажн – θм) = 8,56 a1 = 8,56 |
|  | А a0 = √εвлажн - (a1 * θвлажн) = 1,73 a0 = 1,73 |

Стр. 35

Предметный указатель

Точность, 21

бур, 5, 9

Калибровка

проверьте, 20

обобщённый, 19

сенсор, 15, 18, 19

почва, 15, 16, 19

специфическая почва, 30

Меры предосторожности, 6

Сертификация

емс, 24

Соединения, 10

Преобразования

Линеаризационная таблица, 13, 17, 18, 19

нелинейный, 13, 17, 18

Регистратор данных, 13, 19, 20

DL2e, 12, 19, 20

DL6, 11, 20

GP1, 10, 20

другие, 13

Определения, 25

Диэлектрические свойства, 18

Коэффициент рефракции, 15, 16, 30

Удлинители, 8

GP1, 10, 20

HH2, 5, 14, 15, 20

Установка

Устанавливаемый в почву, 9

Вводной штырь, 5, 8, 9

поверхность, 9

Измеритель, 20

Содержание влаги, 4, 7, 27, 30

Патент, 2

Диэлектрическая проницаемость, 7, 15, 27

Диапазон, 21

Ссылки, 27

Штыри, 4, 6, 7, 9

Код почвы, 5

Солёность, 21, 26

Обслуживание, 29

Почва

глинистая, 16, 17, 25, 30, 31, 33

минеральная, 15, 17, 19, 31, 32, 34

органическая, 15, 17, 19, 25, 32

каменистая, 9

Спецификации, 2, 21

Техническая поддержка, 28, 29

Выявление и устранение неполадок, 20

Проводка, 10

Стр. 36

User Manual for the

SM200

Soil Moisture Sensor

SM200-UM-1.1

Delta-T Devices Ltd

___________________________________________________________________________

Notices

Copyright

All parts of the SM200 design and documentation are the exclusive

right of Delta-T Devices and covered under copyright law, © 2006.

Patent(s) pending.

CE conformity

The SM200 product is CE compliant, conforming to EN61326:1997.

Design changes

Delta-T Devices Ltd reserves the right to change the designs and

specifications of its products at any time without prior notice.

User Manual Version: SM200-UM-1.1 May 2006

Delta-T Devices Ltd

128 Low Road, Burwell

Cambridge CB5 0EJ

UK

Tel: +44 1638 742922

Fax: +44 1638 743155

email: sales@delta-t.co.uk

web: www.delta-t.co.uk

___________________________________________________________________________

Contents

Introduction 4

Description 4

Parts list 5

Care and Safety 6

How the SM200 works 7

Operation 8

Cable attachment 8

Installation 9

Logger connections and configuration 10

GP1 10

DL6 11

DL2e 12

Other data loggers 13

HH2 moisture meter 14

Calibration 15

Soil calibration 16

Sensor calibration 18

Soil moisture reading 19

Troubleshooting 20

Technical Reference 21

Specifications 21

Definitions 25

References 27

Technical Support 28

Soil-specific Calibration 30

Laboratory calibration for non-clay soils 31

Laboratory calibration for clay soils 33

Index 36

SM200 User Manual 1.1 Operation 􀁺 3

Introduction

Description

The SM200 measures soil moisture content at a single location, sampling

~0.5 litres of soil (or other growing medium). It consists of a sealed plastic

body, ~40mm diameter, attached to two sensing rods. The probe is fitted

with a waterproof connector that is easily connected to extension cables.

When taking a reading, the probe rods are inserted directly into the soil.

The output from an SM200 is a simple

analogue dc voltage. This output can be

converted into soil moisture using the

supplied general soil calibrations or the

probe can be calibrated for specific soils.

Features

􀂄 ± 3% scientific accuracy

􀂄 Low salinity sensitivity

􀂄 Excellent temperature stability

􀂄 Minimal soil disturbance

􀂄 Easy installation at depth in augered

holes

􀂄 IP68 connector

􀂄 Rugged, weatherproof and buriable

SM200 User Manual 1.1 Index 􀁺 4

Parts list

You consignment may have the following parts:

| Part | Sales Code | Description |
| --- | --- | --- |
|  | SM200 | SM200 soil moisture sensor fitted with M8 connector, in protective packing tube. |
|  | - (supplied with SM200) | Quick Start Guide |
|  | SM2C/d- HH2 | 1.1m cable for connecting SM200 to HH2 using 25-way D-connector |
|  | SM2C/w-05 | 5m cable connects SM200 to data loggers. Terminates in bare wires. |
|  | EXT/M8-05 EXT/M8-10 EXT/M8-25 | 5, 10 and 25m extension cables. Terminates in IP68 inline connector. |
|  | SM-INRD1 | 1m insertion rod |
|  | ML2-AG1- SL | 45mm spiral auger 1.2m long |

SM200 User Manual 1.1 Introduction 􀁺 5

Care and Safety

| 􀂄 The rods of the SM200 are sharp in order to ease insertion. Care must be taken and handling precautions followed. | To prevent personal injury and damage to the probe always store and transport the SM200 in this protective tube CAUTION SHARP PINS Handle with care |
| --- | --- |
| 􀂄 Avoid touching the rods or exposing them to other sources of static damage. Keep the SM200 in its protective tube when not in use. |  |

􀂄 Take care when attaching cables to ensure that the connectors are

clean, undamaged and properly aligned before pushing the parts

together.

􀂄 Screw together firmly to ensure the connection is water-tight. If the

connector is damaged, apply a light smearing of silicon grease to

maintain sealing.

􀂄 Do not pull the SM200 out of the soil by its cable.

􀂄 If you feel strong resistance when inserting the SM200 into soil, it is

likely you have encountered a stone. Stop pushing and re-insert at a

new location.

SM200 User Manual 1.1 Index 􀁺 6

How the SM200 works

|  | When power is applied to the SM200... |
| --- | --- |
|  | ...it creates a 100MHz waveform (similar to FM radio). |
|  | The waveform is applied to a pair of stainless steel rods which transmit an electromagnetic field into the soil. |
|  | The water content of the soil surrounding the rods... |
|  | ...dominates its permittivity. (A measure of a material’s response to polarisation in an electromagnetic field. Water has a permittivity ≈ 81, compared to soil ≈ 4 and air ≈ 1) |
|  | The permittivity of the soil has a strong influence on the applied field… |
| Vout | …which is detected by the SM200, resulting in a stable voltage output that… |
| Soil Moisture 22 % | ...acts as a simple, sensitive measure of soil moisture content. |

SM200 User Manual 1.1 Operation 􀁺 7

Operation

Cable attachment

| HH2 cable | Extension cables | Logger cable |
| --- | --- | --- |
| 1.1m | 5m 10m 25m | 5m |

􀂄 If you plan to install the SM200 at depth using the insertion rod, pass

the connector through the rod before attaching the cable.

􀂄 Take care when attaching cables to ensure that the connectors are

clean, undamaged and properly aligned before pushing the parts

together.

􀂄 Screw together firmly to ensure the connection is water-tight.

􀂄 Extension cables can be joined up to a recommended maximum of

55m.

Cable functions

The SM200 is a single-ended sensor. The cables are assigned as follows:

| Function | Name | Colour |
| --- | --- | --- |
| Power to sensor | Power V+ | Red |
| Sensor output + | Signal | White |
| Power V- Sensor output – | Ground | Shield (silver) |

SM200 User Manual 1.1 Operation 􀁺 8

Installation

Surface installation and spot

measurements

􀂄 Clear away any stones. Pre-form

holes in very hard soils before

insertion.

􀂄 Push the SM200 into the soil until

the rods are fully inserted. Ensure

good soil contact.

􀂄 If you feel strong resistance when

inserting the SM200, you have probably hit a stone. Stop, and

re-insert at a new location.

Installing at depth

􀂄 Auger a 45mm diameter hole.

~10° to vertical is recommended.

􀂄 Fit insertion rod to SM200 -

remember to pass the cable

through the insertion rod and fit

the connector first.

􀂄 Push the SM200 into the soil until

rods are fully inserted. Ensure

good soil contact.

􀂄 Backfill around the

insertion rod.

Alternatively

􀂄 Dig a trench, and

install horizontally (see

diagram).

SM200 User Manual 1.1 Operation 􀁺 9

Logger connections and configuration

GP1

2 SM200s can be connected to

each GP1, as shown in the

photo.

The diagram and table illustrate

the connections for Channel 1:

| SM200 wiring | Colour | GP1 terminal |
| --- | --- | --- |
| Power V+ | Red | CH1 (PWR) |
| Signal | White | CH1 (+) |
| Ground | Shield | CH1 (GND) and CH1 (-) |

Configure the GP1 by choosing “SM200” from the sensor menu in

DeltaLINK-PC or Pocket DeltaLINK.

For configuration details see the GP1 Quick Start Guide.

SM200 User Manual 1.1 Index 􀁺 10

DL6

Up to 6 SM200s can be connected

to each DL6.

These details illustrate connection

to Channel1:

| SM200 wiring | Colour | DL6 terminal |
| --- | --- | --- |
| Power V+ | Red | V+ |
| Signal | White | IN+ |
| Ground | Shield | IN- and 0V |

Configure the DL6 by choosing “SM200” from the sensor menu in

DeltaLINK-PC or Pocket DeltaLINK. For configuration details

refer to the DL6 Quick Start Guide.

SM200 User Manual 1.1 Operation 􀁺 11

DL2e

Configure the DL2e by choosing one of the SM200 sensor codes from

the LS2Win sensor library, according to your soil type.

Up to 60 SM200s can be connected

to each DL2e.

These details illustrate connection

to Channel 58 using the LAC1 input

card configured in 15-channel

mode, and warm-up channel 63:

| SM200 wiring | Colour | DL2e terminal |
| --- | --- | --- |
| Power V+ | Red | 63 NO |
| Signal | White | 58+ |
| Ground | Shield | 58- and 61- |

For configuration details see the DL2e User Documentation.

SM200 User Manual 1.1 Index 􀁺 12

Other data loggers

􀂄 The SM200 should be connected as a single-ended, powered

sensor.

􀂄 Configure the logger input as a voltage sensor, approximately

0 to 1.0 volts corresponding to ~0 to 60% water content – see

next section.

􀂄 The SM200 has been optimised for a 1 second warm-up

period. It is recommended that the sensor is not powered

continuously.

􀂄 Configure the logger to convert the SM200 readings into soil

moisture using the polynomial or linearisation tables – see

Calibration.

SM200 User Manual 1.1 Operation 􀁺 13

HH2 moisture meter

􀂄 Connect the SM200 to the HH2 meter.

􀂄 Press Esc to turn the meter on, and if

necessary press again until the HH2

displays the start-up screen.

􀂄 Set the meter to read from an SM200:

► Press Set and scroll down to the

Device option.

► Press Set again and scroll down

to select SM200:

Device:

SM200

► Press Set to confirm this choice.

􀂄 If you intend to store readings, it may be helpful to define each

reading by setting a plot label and sample number – accessed

by pressing Set and scrolling to the Data option.

􀂄 Press Read to take a reading (it takes about 2 seconds).

You can choose different units from the Display option.

SM200 Store?

20.3%vol

􀂄 Press Store to save or Esc to discard the reading.

􀂄 Remove the SM200 and move to a new location...

If you have saved data, connect your HH2 to a PC and run

HH2Read to retrieve the readings.

Refer to the HH2 User Manual for further details.

SM200 User Manual 1.1 Index 􀁺 14

Calibration

The SM200 is provided with general calibrations for mineral and

organic soils which can be used to convert the output from the

sensor directly into soil moisture when used with Delta-T loggers

and the HH2 moisture meter. This section explains how these

calibrations work, how to adapt them for other soils and how to

provide calibrations for other data loggers.

The SM200 measures volumetric soil moisture, θ, by detecting the

dielectric properties of the damp soil – the permittivity, ε, or more

conveniently the refractive index, which is closely equivalent to

√ε. The SM200 response is best understood in these stages:

| 1. Soil calibration | + | 2. Sensor calibration |
| --- | --- | --- |
| θ → √ε |  | V → √ε |
| Soil calibrations График √ε (θ) |  | SM200 dielectric performance График √ε (V) |

3. Soil moisture reading

V → θ

SM200 soil moisture reading

График θ (V)

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 15

Soil calibration

Damp soil is essentially a mixture of soil particles, air and water,

and together these components determine its dielectric properties,

including the refractive index √ε. The refractive index of the

mixture is approximated simply by adding the contributions from the

individual components [ref 4.].

For a particular soil, the contribution from the soil particles can be

assumed to be constant, so the refractive index measured by the

SM200 is only affected by changes to the water content, θ. This

relationship simplifies to:

√ε = a0+a1* θ

where the coefficients a0 and a1 conveniently parameterise the

dielectric properties of soils.

График Soil calibrations

Ось X - Soil moisture (m3.m-3)

Ось Y - Soil refractive index (√ε)

Offset (a0)

Slope (a1)

Note that:

a0 = √εdry_soil is usually between 1.3 to 2.3

a1 corresponds approximately to √εwater – 1 and usually takes a

value about 8.0. Real soil values for a0 and a1 can vary

significantly from these guidelines when they are affected by other

factors – in particular, bound water in clay soils may result in higher

values of a1.

SM200 User Manual 1.1 Index 􀁺 16

General soil calibrations

Most soils can be characterised simply by choosing one of the two

general calibrations we supply, one for mineral soils (predominantly

sand, silt and clay) and one for organic soils (with a high organic

matter content).

|  | a0 | a1 |
| --- | --- | --- |
| Mineral soils | 1.6 | 8.4 |
| Organic soils | 1.3 | 7.7 |

График  General soil calibrations

Ось X - Soil Moisture (m3.m-3)

Ось Y - Soil refractive index (√ε)

Red – mineral

Green - organic

These values have been used to generate the polynomial

conversions and linearisation tables in the Soil moisture reading

section.

Soil-specific calibration

Instead of adopting these general calibrations, you may wish to

determine specific calibration values of a0 and a1 for your soil. This

procedure is fairly straightforward if you can get access to standard

laboratory equipment and is described in detail in Appendix A.

Soil specific calibration can significantly improve the accuracy of

individual readings - but make less of an improvement to readings

where installation and sampling errors are high.

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 17

Sensor calibration

Each SM200 is individually adjusted to provide consistent dielectric

performance:

График  SM200 dielectric performance

Ось X - SM200 output (Volts)

Ось Y - Soil refractive index (√ε)

This response can be approximated either by a polynomial or by a

linearisation table:

Polynomial (use for SM200 readings from 0 to 1.1 Volts)

√ε = 1,0+16,103V – 38,725V2 + 60,881V3 – 46,032V4 + 13,536V5

where V is the SM200 output in Volts

Linearisation table (use for SM200 readings from 0 to 1.2 Volts)

| V | √ε | V | √ε | V | √ε | V | √ε |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0.000 | 1.011 | 0.150 | 2.729 | 0.400 | 4.102 | 1.100 | 7.290 |
| 0.030 | 1.453 | 0.180 | 2.956 | 0.600 | 4.956 | 1.150 | 7.610 |
| 0.060 | 1.842 | 0.220 | 3.218 | 0.820 | 5.939 | 1.200 | 7.982 |
| 0.090 | 2.179 | 0.260 | 3.446 | 0.960 | 6.571 | 1.220 | 8.205 |
| 0.120 | 2.473 | 0.310 | 3.701 | 1.040 | 6.960 | 1.800 | 9.000 |

SM200 User Manual 1.1 Index 􀁺 18

Soil moisture reading

Polynomial conversion

Combining the Soil calibrations and Sensor calibration

steps, the conversion equation becomes:

θ=

where a0 and a1 are the calibration coefficients.

For a generalised mineral soil this becomes:

And for a generalised organic soil:

Linearisation table conversion

Alternatively, the conversion from SM200 reading (Volts) to soil

moisture θ (m3.m-3 or %vol) can be accomplished by a linearisation

table.

The following table lists the values used for the DL2e data logger:

| Soil moisture | Mineral soil | Organic soil | Soil moisture | Mineral soil | Organic soil |
| --- | --- | --- | --- | --- | --- |
| %vol | Volts | Volts | %vol | Volts | Volts |
| -4 | -2.090 | -2.090 | 52 | 0.826 | 0.678 |
| 0 | 0.041 | 0.019 | 56 | 0.901 | 0.747 |
| 4 | 0.068 | 0.041 | 60 | 0.975 | 0.816 |
| 8 | 0.099 | 0.066 | 64 | 1.043 | 0.884 |
| 12 | 0.135 | 0.094 | 68 | 1.103 | 0.952 |
| 16 | 0.178 | 0.126 | 72 | 1.156 | 1.017 |
| 20 | 0.230 | 0.164 | 76 | 1.199 | 1.076 |
| 24 | 0.292 | 0.209 | 80 | 1.237 | 1.127 |
| 28 | 0.365 | 0.262 | 84 | 1.269 | 1.171 |
| 32 | 0.443 | 0.324 | 88 | 1.297 | 1.210 |
| 36 | 0.523 | 0.393 | 92 | 1.322 | 1.244 |
| 40 | 0.601 | 0.466 | 96 | 1.347 | 1.278 |
| 44 | 0.676 | 0.538 | 100 | 1.800 | 1.800 |
| 48 | 0.751 | 0.609 | 104 | 2.090 | 2.090 |

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 19

Troubleshooting

Always try to identify which part of the measurement system is the

source of the difficulty. For the SM200 this may fall into one of the

following areas:

The measurement device

What equipment is being used to read the probe output?

• A Delta-T HH2 Moisture Meter

• A Delta-T data logger such as the GP1, DL6 or DL2e

Consult the user manuals or the on-line help for these devices and

their related software. Try alternative types of equipment if you

have them available.

Check that you are using an appropriate soil calibration and the

correct conversion method – see Calibration section.

The SM200 itself

Try to isolate the problem into one of the following areas

• The SM200 or the connecting cable

Then try to narrow down the area further

• Mechanical problems faults, or damage

• Electrical or electronic problems or faults

Functional check

The following two simple checks can be used to establish whether

your SM200 is functioning within expected bounds:

Air reading

Hold the SM200 away from other objects and take a reading using

an HH2 meter, or other meter or logger. The reading should be 0

±12mV when used with a 5m cable.

Finger reading

Hold the SM200 with your fingers wrapped closely (but carefully!)

around the two metal rods and take a reading using an HH2 meter

or other meter or logger. The reading should be between 500 to

750mV.

SM200 User Manual 1.1 Index 􀁺 20

Technical Reference

Specifications

|  |  | Range and Notes |  |
| --- | --- | --- | --- |
| Accuracy* | ± 3 %vol | 0 to 50 %vol |  |
| Salinity errors | ± 3.5 %vol | 0 to 40 %vol, over | 50 to 500 mS.m-1 (0.5 to 5 dS.m-1) |
| Power requirement | 5 to 14V 15mA typical | 1 second warm-up recommended 200ms minimum |  |
| Output signal | 0 to 1.0V | Nominal 0 to 60 %vol |  |
| Environmental sealing | IP68 | Sensor buriable, sealed to IP68 |  |
| Operating range | -20 to +60°C | Readings in frozen soil are not meaningful |  |

* Accuracy figures apply to calibration soils and over 0 to 60°C.

Percentages quoted as %vol are % volumetric water content.

Conductivity response

График SM200 conductivity response at different water contents

Ось X Conductivity ECp (mS/m)

Ось Y SM200 output (Volts)

100%(вода)

Розовая кривая: Typical artificial substrate conditions

Зелёная кривая: Typical soil conditions

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 21

Temperature response

The SM200 temperature response in any particular soil will depend on a

combination of effects:

􀂄 The SM200 electronics has a very low temperature sensitivity, and

makes a negligible contribution to the overall sensitivity.

􀂄 The refractive index of water (√ε, see Calibration section) reduces as

the temperature increases. This produces a negative temperature

response particularly in soils or substrates with high water content.

􀂄 Water that is bound to the surface of soil particles has a much lower

refractive index than free water. The percentage of bound water

decreases with temperature and this produces a positive temperature

response particularly in clay soils at lower water contents.

The last two effects partially offset each other, but in soil conditions where

one or the other effect dominates, the SM200 will appear to have a

significant temperature response. This illustration is based on the model in

reference 7:

График Illustrating temperature dependence of SM200 readings in clay soil

Ось X Soil temperature (°C)

Ось Y Soil moisture content (m3.m-3)

Розовая линия – bound water

Note: ice has a quite different refractive index from water, so SM200

readings cannot be interpreted reliably when inserted into soil below 0°C.

SM200 User Manual 1.1 Index 􀁺 22

Sampling Volume

The SM200 is most sensitive to signals very close to the two rods, but a

small proportion of the signal reaches up to 50mm from the rods. For best

accuracy we recommend taking readings in at least 0.5 litres of soil,

preferably 1.0 litre.

SM200s may interact if they are placed too close together – they should be

separated by at least 100mm.

The sensitivity field for an SM200 inserted in water has been tested as

shown in this figure (the vertical scale is in arbitrary units):

График SM200 field of sensitivity surrounding the rods

Ось X Distance parallel to plane of rods from a centre line between yhe rods (mm)

Ось Z Distance perpendicular to plane of rods (mm)

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 23

Electromagnetic Compatibility (EMC)

General information

SM200 is a Class A product, intended for operation in nonresidential

environments.

Only use cables and accessories authorised by Delta-T (sensor

cables from other sources for example can adversely affect product

performance and affect quality of results).

If possible route cables along the soil surface or bury them – this

also reduces possible trip hazard and animal damage.

Do not modify the product or its supplied accessories.

Regulatory information

Europe

SM200 has been assessed for CE compliance under the European

Union EMC directive 89/336/EEC and conforms to the appropriate

standards, EN61326 (1997), providing the product is used as

described in this manual. Use of non Delta-T cables and

accessories may invalidate this.

FCC compliance (USA)

This device complies with part 15 of the FCC Rules. Operation is

subject to the following conditions:

1. This device may not cause harmful interference

2. This device must accept any interference received, including

interference that may cause undesired operation

Note: This equipment has been tested and found to comply with the

limits for a Class A digital device, pursuant to part 15 of the FCC

rules. These limits are designed to provide reasonable protection

against harmful interference when the equipment is operated in a

commercial environment. This equipment generates, uses, and can

radiate radio frequency energy and, if not installed and used in

accordance with the instruction manual, may cause harmful

interference to radio communications. Operation of this equipment

in a residential area is likely to cause harmful interference in which

case the user will be required to correct the interference at his/her

own expense.

Modifications not expressly approved by the manufacturer could

void the user’s authority to operate the equipment under FCC rules

SM200 User Manual 1.1 Index 􀁺 24

Definitions

Volumetric Soil Moisture Content is defined as

Θv= Vw/Vs

where Vw is the volume of water contained in the sample

and Vs is the total volume of the soil sample.

The preferred units for this ratio are m3.m-3, though %vol is

frequently used.

Soil Moisture Content varies from approx. 0.02 m3.m-3 for sandy

soils at the permanent wilting point, through approx. 0.4 m3.m-3 for

clay soils at their field capacity, up to values as high as 0.85 m3.m-3

in saturated peat soils.

Gravimetric Soil Moisture Content is defined as

Θg= (Mw/Ms)*g*g-1

where MW is the mass of water in the sample,

and MS is the total mass of the dry sample.

To convert from volumetric to gravimetric water content, use the

equation

Θg= Θv *(ρw/ρs)

where ρw is the density of water (= 1g.cm-3),

and ρS is the bulk density of the sample (Ms/Vs)

Organic and Mineral soil definitions:

The general calibrations have been optimised to cover a wide

range of soil types, based on the following definitions:

| Soil type | optimised around organic content: | use for organic contents: | bulk density range: (g.cm-3) | use for bulk densities: (g.cm-3) |
| --- | --- | --- | --- | --- |
| Mineral | ~ 1 %C | < 7 %C | 1.25 - 1.5 | > 1.0 |
| Organic | ~ 40 %C | > 7 %C | 0.2 - 0.7 | < 1.0 |

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 25

Salinity

The preferred SI units for ionic conductivity are mS.m-1 (where S is

Siemens, the unit of electric conductance = ohm-1).

The following conversions apply:

1 mS.m-1   = 0.01 dS.m-1

= 0.01 mS.cm-1

= 10 µS.cm-1

Soil salinity is also partitioned into the following descriptive

categories:

non-saline                 0 - 200                      mS.m-1

slightly saline            200 - 400                  mS.m-1

moderately saline     400 - 800                  mS.m-1

strongly saline           800 - 1600               mS.m-1

extremely saline        > 1600                     mS.m-1

SM200 User Manual 1.1 Index 􀁺 26

References

1. Gaskin, G.J. and J.D. Miller, 1996

Measurement of soil water content using a simplified impedance

measuring technique.

J. Agr. Engng Res 63, 153-160

2. Topp, G.C., J. L. Davis and A. P Annan 1980

Electromagnetic determination of soil water content.

Water Resour. Res 16(3) 574-582

3. Whalley, W.R. 1993

Considerations on the use of time-domain reflectometry (TDR) for

measuring soil moisture content.

Journal of Soil Sci. 44, 1-9

4. White, I., J.H. Knight, S.J. Zegelin, and Topp, G.C. 1994

Comments on ‘Considerations on the use of time-domain

reflectometry (TDR) for measuring soil water content’ by W R

Whalley

Journal of Soil Sci. 45, 503-508

5. Roth, C.H., M.A. Malicki, and R. Plagge, 1992

Empirical evaluation of the relationship between soil dielectric

constant and volumetric water content as the basis for calibrating

soil moisture measurements.

Journal of Soil Sci. 43, 1-13

6. Knight, J.H. 1992

Sensitivity of Time Domain Reflectometry measurements to lateral

variations in soil water content.

Water Resour. Res., 28, 2345-2352

7. Or, D. and J.M. Wraith 1999

Temperature effects on soil bulk dielectric permittivity measured by

time domain reflectrometry: A physical model.

Water Resour Res., 35, 371-383

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 27

Technical Support

Terms and Conditions of Sale

Our Conditions of Sale (ref: COND: 1/00) set out Delta-T's legal obligations

on these matters. The following paragraphs summarise Delta-T's position

but reference should always be made to the exact terms of our Conditions

of Sale, which will prevail over the following explanation.

Delta-T warrants that the goods will be free from defects arising out of the

materials used or poor workmanship for a period of twelve months from

the date of delivery.

Delta-T shall be under no liability in respect of any defect arising from fair

wear and tear and the warranty does not cover damage through misuse or

inexpert servicing, or other circumstances beyond our control.

If the buyer experiences problems with the goods they shall notify Delta-T

(or Delta-T’s local distributor) as soon as they become aware of such

problem.

Delta-T may rectify the problem by supplying replacement parts free of

charge, or by repairing the goods free of charge at Delta-T's premises in

the UK, during the warranty period.

If Delta-T requires that goods under warranty be returned to them from

overseas for repair, Delta-T shall not be liable for the cost of carriage or for

customs clearance in respect of such goods. However, we much prefer to

have such returns discussed with us in advance and we may, at our

discretion, waive these charges.

Delta-T shall not be liable to supply products free of charge or repair any

goods where the products or goods in question have been discontinued or

have become obsolete, although Delta-T will endeavour to remedy the

buyer’s problem.

Delta-T shall not be liable to the buyer for any consequential loss, damage

or compensation whatsoever (whether caused by the negligence of the

Delta-T, our employees or distributors or otherwise) which arise from the

supply of the goods and/or services, or their use or resale by the buyer.

Delta-T shall not be liable to the buyer by reason of any delay or failure to

perform our obligations in relation to the goods and/or services, if the delay

or failure was due to any cause beyond the Delta-T’s reasonable control.

SM200 User Manual 1.1 Index 􀁺 28

Service and Spares

Users in countries that have a Delta-T Distributor or Technical

Representative should contact them in the first instance.

Spare parts for our own instruments can be supplied from our works. These

can normally be despatched within a few working days of receiving an

order.

Spare parts and accessories for sensors or other products not

manufactured by Delta-T may have to be obtained from our supplier and a

certain amount of additional delay is inevitable.

No goods or equipment should be returned to Delta-T without first obtaining

the agreement of Delta-T or our distributor.

On receipt at Delta-T, the goods will be inspected and the user informed of

the likely cost and delay. We normally expect to complete repairs within a

few working days of receiving the equipment. However, if the equipment

has to be forwarded to our original supplier for specialist repairs or

recalibration, additional delays of a few weeks may be expected.

Technical Support

Technical Support is available on Delta-T products and systems. Users in

countries that have a Delta-T Distributor or Technical Representative

should contact them in the first instance.

Technical Support questions received by Delta-T will be handled by our

Tech Support team. Your initial enquiry will be acknowledged immediately

with a “T number” and an estimate of time for a detailed reply. Make sure to

quote our T number subsequently so that we can easily trace any earlier

correspondence.

In your enquiry, always quote instrument serial numbers, software version

numbers, and the approximate date and source of purchase where these

are relevant.

Contact details:

Tech Support Team

Delta-T Devices Ltd

128 Low Road, Burwell, Cambridge CB5 0EJ, U.K.

email: tech.support@delta-t.co.uk

Web site: www.delta-t.co.uk

Tel: +44 (0) 1638 742922

Fax: +44 (0) 1638 743155

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 29

Soil-specific Calibration

This note provides details of 2 techniques for generating soil-specific

calibrations:

1. Laboratory calibration for substrates* and non-clay soils

2. Laboratory calibration for clay soils

* We use the term substrate to refer to any artificial growing medium.

Underlying principle

Soil moisture content (θ) is proportional to the refractive index of the soil

(√ε) as measured by the SM200 (see Calibration section).

График Dielectric property of soil compared to moisture content

Ось X: moisture content (m3*m-3)

Ось Y: √ε

The goal of calibration is to generate two coefficients (a0, a1) which can be

used in a linear equation to convert probe readings into soil moisture:

√ε = a0 + a1 * θ

SM200 User Manual 1.1 Index 􀁺 30

Laboratory calibration for non-clay soils

This is the easiest technique, but it’s not suitable for soils that shrink or

become very hard when dry.

Equipment you will need:

􀂄 SM200 and meter

􀂄 Soil corer (if doing a calibration for a cohesive soil rather than sand or a

substrate)

􀂄 Heat-resistant beaker (≥ 0.5 litre)

􀂄 Weighing balance (accurate to < 1g)

􀂄 Temperature controlled oven (for mineral soils or substrates)

| Process | Notes and example |
| --- | --- |
|  | Collect a damp sample of the soil or substrate. This sample needs to be unchanged from its in-situ density, to be ≥ 0.5 litre, to have the correct dimensions to fit the beaker, and to be generally uniform in water content. For cohesive soils this is most easily done with a soil-corer. Sandy soils can be poured into the beaker, but you should take the subsequent measurements immediately, as the water will quickly begin to drain to the bottom of the beaker. Compressible soils and composts often require measurement of the in-situ density and then need to be carefully reconstituted at that density within the beaker. |
|  | Measure the volume occupied by the sample. Ls = 463.5ml |
|  | Weigh the sample, including the beaker. Ww = 743.3g |

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 31

|  | Insert SM200 into the sample and record its output in Volts. Vw = 0.350V |
| --- | --- |
|  | Dry the sample thoroughly. With mineral soils this is usually achieved by keeping it in the oven at 105°C for several hours or days (the time required depends on the sample size and porosity). For organic soils and composts it’s usual to air-dry the sample to avoid burning off any volatile factions. |
|  | Weigh the dry sample in the beaker. W0 = 627.2g |
|  | Re-insert the SM200 into the dry sample and record this reading. V0 = 0.051V |
| Calculate a0 | For the SM200, √ε = 1,0+16,103V – 38,725V2 + 60,881V3 – 46,032V4 + 13,536V5 In the dry soil V = V0 = 0.051 Volts, and substituting this value into the above equation gives √ε0 = 1,73 Since θ 0 = 0, this is the value needed for a0 a0 = 1.73 |
| Calculate θw | The water content of the wet soil, θw, can be calculated from the weight of water lost during drying, (Ww – W0) and its volume, Ls: θv = (Ww – W0)/Ls = (743,3 – 627,2)/463,5 = 0,25 θv = 0,25 |

SM200 User Manual 1.1 Index 􀁺 32

| Calculate a1 | In the wet soil V = Vw = 0.350 Volts and substituting gives √εw = 3,88 Finally a1 = (√εw - √ε0)/(θw - θ0) = (3,88 – 1,73)/(0,25 - 0) = 8,60 a1 = 8.60 |
| --- | --- |

Laboratory calibration for clay soils

This technique is adapted to avoid the near-impossibility of inserting the

SM200 into completely dry clay soil. It requires taking measurements at 2

significantly different, but still damp, moisture levels.

Equipment you will need:

􀂄 SM200 and meter

􀂄 Soil corer

􀂄 Heat-resistant beaker (≥ 500ml)

􀂄 Weighing balance (accurate to < 1g)

􀂄 Temperature controlled oven

| Process | Notes and example |
| --- | --- |
|  | Collect a wet sample of the clay soil: 25 to 30% water content would be ideal. This sample needs to be unchanged from its in-situ density, to be ≥ 500ml, to have the correct dimensions to fit the beaker, and to be generally uniform in water content. This is most easily done with soil-corer. |
|  | Measure the volume occupied by the sample. Ls = 463.5ml |
|  | Weigh the wet sample, including the beaker. Ww = 743.3g |

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 33

|  | Insert SM200 into the wet sample and record its output in Volts. Vw = 0.349V |
| --- | --- |
|  | Dry the sample until still moist, ~15% water content. Gentle warming can be used to accelerate the process, but take care not to over-dry in places, and allow time for the water content to equilibrate throughout the sample before taking a reading. |
|  | Reweigh. Wm = 693.2g |
|  | Re-measure with the SM200. Vm = 0.180V |
|  | Dry the sample thoroughly. With mineral soils this is usually achieved by keeping it in the oven at 105°C for several hours or days (the time required depends on the sample size and porosity). |
|  | Weigh the dry sample in the beaker. W0 = 627.2g |

SM200 User Manual 1.1 Index 􀁺 34

| Calculations | Substituting in the SM200 equation √ε = 1,0+16,103V – 38,725V2 + 60,881V3 – 46,032V4 + 13,536V5 provides two dielectric values, √εw and √εm, at two known water contents, θw and θm |
| --- | --- |
| For the wet soil | Substituting Vw = 0.349 gives √εw = 3,88 = a0 + a1 * θw For θw = (743,3 – 627,2)/463,5 = 0,25 |
| For the moist soil | Substituting Vm = 0.180 gives √εm = 2,95 = a0 + a1 * θm For θm = (693,2 – 627,2)/463,5 = 0,14 |
| Calculate a1 | Then a1 = (√εw - √εm)/(θw – θm) = 8,56 a1 = 8.56 |
| Calculate a0 | And a0 = √εw - (a1 * θw) = 1,73 a0 = 1.73 |

SM200 User Manual 1.1 Soil-specific Calibration 􀁺 35

Index

A

Accuracy, 21

Auger, 5, 9

C

Calibration

check, 20

generalised, 19

sensor, 15, 18, 19

soil, 15, 16, 19

soil-specific, 30

Care and safety, 6

Certification

emc, 24

Connections, 10

Conversions

linearisation table, 13, 17, 18, 19

polynomial, 13, 17, 18

D

Data logger, 13, 19, 20

DL2e, 12, 19, 20

DL6, 11, 20

GP1, 10, 20

other, 13

Definitions, 25

Dielectric

performance, 18

refractive index, 15, 16, 30

E

Extension cables, 8

G

GP1, 10, 20

H

HH2, 5, 14, 15, 20

I

Installation

buried, 9

insertion rod, 5, 8, 9

surface, 9

M

Meter, 20

Moisture content, 4, 7, 27, 30

P

Patent, 2

Permittivity, 7, 15, 27

R

Range, 21

References, 27

Rods, 4, 6, 7, 9

S

Sales code, 5

Salinity, 21, 26

Servicing, 29

Soil

clay, 16, 17, 25, 30, 31, 33

mineral, 15, 17, 19, 31, 32, 34

organic, 15, 17, 19, 25, 32

stony, 9

Specifications, 2, 21

T

Technical support, 28, 29

Troubleshooting, 20

W

Wiring, 10

SM200 User Manual 1.1 Index 􀁺 36